Response to Supervisor’s Comments

ECG educational open-hardware manuscript

Thank you for the detailed review. The manuscript has been repositioned from a preliminary biomedical-device report to an Educational Hardware article. The revision does not invent unavailable measurements; instead, it narrows the claims to the assembled prototype, retained functional evidence and a complete open educational release.

# Point-by-point response

| Comment | Action taken |
| --- | --- |
| Reason for choosing the journal | The project’s strongest contribution is an openly documented teaching platform rather than a clinically validated ECG device. The Journal of Open Hardware’s Educational Hardware category is therefore the closest scope match. |
| Open-source release | A repository-ready package named `open-educational-ecg-data-logger` has been completed with editable SVG wiring, reference firmware, BOM, assembly guide, evidence matrix, test protocol, safety records, media and licence information. Proposed review URL: https://github.com/gastonravindias/open-educational-ecg-data-logger; release v1.0.0. |
| Repository wording | Future-tense descriptions of intended files were removed. The manuscript now describes the completed package. The only external action is activating the public GitHub repository before submission. |
| Solar-panel calculation | Because the original panel label and power records were unavailable, no measured solar claim is made. A transparent replication design calculation is provided: a conservative 0.60 W load and stated efficiency assumptions give 2.79 W minimum, leading to a ≥3 W recommendation. |
| Reproducibility | The manuscript states that one original build was completed and that independent replication was not performed. It does not claim a non-existent second build. The release supplies the documentation needed for future replication. |
| Quantitative quality control | Unavailable timing, dropped-sample, SD-reliability, runtime and charging results are explicitly marked as not evaluated. The released firmware provides a nominal 250 samples/s schedule, raw 12-bit ADC logging and missed-interval tracking, while a test protocol is supplied for future measurements. |
| Safety engineering | Battery-only electrode operation, physical disconnection of USB/charging equipment, no human-connected use of the open breadboard, protected power components and simulator-based quantitative testing are now explicit engineering controls. |
| Title and abstract | The title was shortened and solar performance was removed from the central claim. The abstract now reports the available evidence, code-defined settings and limitations. |
| System architecture | A concise signal-chain paragraph and a revised architecture figure now show electrodes → AD8232 → ESP32-C3 → display/RTC/microSD, with separate power telemetry. |
| IQPS terminology | The term ‘Intelligent Quality-Power Switching’ was removed. The release uses deterministic status logic and explicitly states that no trained AI or diagnostic model is implemented. |
| Application | Educational use is separated from non-clinical engineering reuse. Diagnosis, treatment, patient monitoring and unsupervised human use are excluded. |
| References and licences | Background references were trimmed and open-hardware and safety-oriented references were added. The release allocates CC BY 4.0 to the article/documentation, CERN-OHL-P-2.0 to hardware files and MIT to firmware. |

# Remaining submission action

Create the public GitHub repository at https://github.com/gastonravindias/open-educational-ecg-data-logger, upload the contents of the supplied ZIP, and create release v1.0.0. Then use the manuscript named `READY_AFTER_GITHUB_UPLOAD`. No additional experimental data are inserted or claimed.